Article title
A General Purpose Alarm Device (GPAD)

Authors
Robert L. Read*, Lawrence Kincheloe, Forrest Lee Erickson

Affiliations
Robert L. Read	Lawrence R. Kincheloe III	Forrest Lee Erickson
Public Invention	SPEC & Public Invention	Public Invention
1709 Norris Dr.	3217 Brush Creek Road	1709 Norris Dr.
Austin, TX, 78704	Oklahoma City, OK 73120	Austin, TX, 78704

Corresponding author’s email address and Twitter handle
read.robert@pubinv.org
twitter: @robertleeread

Abstract

The General Purpose Alarm Device (GPAD) shines lights and makes loud noises to draw the attention of a human being to a problem. It has a programmable 80 character display to provide textual info. A mute button toggles the sound on and off. Fundamentally it is designed to act as a peripheral to a controlling computer or microcontroller. The controller may communicate over a USB (COM) connection or through a 5V SPI connection via an RJ12 cable. The GPAD is intended to be as general as possible, so that it can be used to alarm many engineering, scientific, and hobby machines, instruments, and situations. The original driving use case is to provide medical alarm capability to the PolyVent [5]open-source mechanical medical ventilator. The GPAD supports 5 alarm levels above “silent” of increasing urgency in terms of light, rhythm, and frequency. It is based on the Arduino Uno Atmega328 design and potentially extensible the same way an Uno is, through headers and shields. It includes the printed wiring assembly, firmware for the GPAD peripheral, a simple documented API and a 3D printable enclosure. The repo includes instructions for using a second GPAD as a controller.

Keywords
alarms, arduino, UNO, SPI, medical alarms, monitoring

Specifications table

Hardware name	General Purpose Alarm Device (GPAD) v2.0
Subject area	General
Hardware type	Other: general alarms for digitally detectable alert conditions
Closest commercial analog	No known commercial analog is available.
Open source license	CERN Open Hardware Licence Version 2 - Strongly Reciprocal
Cost of hardware	$100.00
Source file repository	If you’ve uploaded your source files to an approved repository (OSF, Mendeley Data or Zenodo) write the DOI URL here. / For example: http://doi.org/10.17605/OSF.IO/WGK7Q
OSHWA certification UID (OPTIONAL)	US002352

Hardware in context

We are not aware of any alarm device that has the programmable features of the GPAD. Electronic component firms such as Mouser and Digikey sell buzzers, sirens, and loudspeakers which can make noise; they sell lights that can represent an alarm condition; they sell proprietary alarm devices that are tired to special purposes (such as a smoke alarm), but we know of know digitally controllable system that combines these features. The Adafruit Towerlight [6], an example of a “stacklight”, but it does not seem to have a digital interface, and has no text output capability, and no abstract API. Mucco makes a complete line [5] of line of integrated warning horns and stack lights, but they appear to have no digital interface, no text, and no alarm levels; they are simply on or off. Although we believe a GPAD is a widely usable device, its use is of course confined to manufacturers, makers, or engineers who are capable of programming it, which may preclude the market volume associated with consumer devices.

The digital controllers for 3D printers can often make a beep, and have a digital display, which is sometimes backlit. This has some of the features of the GPAD, but is not as bright or as loud, nor programmed for this purpose as conveniently. A cheap example is the HiLetGo 3-01-0889 [7].

Hardware description

We have designed a printed wiring assembly and an enclosure. The enclosure is approximately 160mm x 110mm x 50mm (6”x4.25”x1.75”) The assembly weighs 232 grams (exclusive of any power supply).
The device has a 20x4 character display (English characters), five bright white programmable LEDs, a programmable buzzer, a “Mute” button which silences the buzzer and a non programmable power indicator LED. The enclosure hangs by (for example) tie wraps from a rail for vertical display or can rest on a horizontal surface. Typical signal connection is by an RJ12 data cable. Power is by an external wall supply with a 2.1mm barrel jack power connector at 9-12V at <1000 mA.

The GPAD integrates several features hobbyists find useful in one device. Although you can purchase 20x4 displays with parallel or I2C interfaces it is the integration in one PCB and enclosure of all of the features which our device adds to make this tethered device integrate into a system easily.

The GPAD hardware features the following:
- Five bright LEDs to visually indicate alerts.
- A buzzer to indicate urgent attention is required.
- A button for minimum user feedback and to mute alerts.
- A 20x4 character display to provide detailed information to the user.
- An enclosure that can be vertically mounted by hanging, for example with wire ties.

Scientific experiments and laboratory equipment often monitor something over time whose condition can reach a state that requires human attention. The GPAD can thus be useful in the field or laboratory for:
- Alerting to conditions which may ruin the experiment if not corrected,
- Alerting to the very conditions sought by the experimenter and therefore demanding attention.
- Indicating potential failure of the equipment or the need for maintenance, such as low battery charge, over-temperature, over-pressure, etc.
The GPAD is designed to be loosely integrated with new devices. At the cost of modifying the enclosure, a researcher may integrate the GPAD into their own equipment.

Design files summary

Design file name	File type	Open source license	Location of the file
Firmware	A folder with source files. A README.md there describes the various source files	Firmware: Affero GPL 3.0	https://github.com/PubInv/general-purpose-alarm-device/tree/main/Firmware
Simulation	Wokwi, simulations for Factory Test and other development efforts.	Firmware: Affero GPL 3.0	https://github.com/PubInv/general-purpose-alarm-device/tree/main/simulation
Hardware	Folder holding a README describing the subfolders to follow.	CERN Open Hardware Licence Version 2 - Strongly Reciprocal	https://github.com/PubInv/general-purpose-alarm-device/tree/main/Hardware
GeneralPurposeAlarmDevicePCB	KiCad schematic and PCB files	CERN Open Hardware Licence Version 2 - Strongly Reciprocal	https://github.com/PubInv/general-purpose-alarm-device/tree/main/Hardware/GeneralPurposeAlarmDevicePCB
Enclosure	OpenSCAD and FreeCAD source files, STL and STEP files for the enclosure	CERN Open Hardware Licence Version 2 - Strongly Reciprocal	https://github.com/PubInv/general-purpose-alarm-device/tree/main/Hardware/Enclosure
Manufacturing	Bill of materials, Gerber files and description records of parts ordered for the builds.	CERN Open Hardware Licence Version 2 - Strongly Reciprocal	https://github.com/PubInv/general-purpose-alarm-device/tree/main/Hardware/Manufacturing
Manufacturing and Unit Test Documentation	Assembly instructions and notes	CERN Open Hardware Licence Version 2 - Strongly Reciprocal	https://github.com/PubInv/general-purpose-alarm-device/blob/main/Hardware/Documentation/ManufacturingUnitTestTroubleshootingRev2.md

Bill of materials summary

The spread sheet used to receive the KiCad generated BOM/PARTS LIST and to transform it into the files necessary for ordering PCB assemblies from JLCPCB is found at:
https://github.com/PubInv/general-purpose-alarm-device/blob/main/Hardware/GeneralPurposeAlarmDevicePCB/GPAD_V2_BOM.xls

Instructions are built into the spreadsheet for its use. JLCPCB built the raw PCB and did SMT assembly and some through hole assembly. The balance of the through hole assembly was done by Public Invention volunteers.

Designator	Component	Number	Cost per unit -currency	Total cost - / currency	Source of materials	Material type
Enclosure LED Standoff	U_Box_V105_GPAD_LED_Standoff_single.stl	6	$1.01	$1.01	Fabricated by JLCPCB	321PA-F NYLON
Enclosure Frontpanel	U_Box_V105_General_Alarm_Device_FrontPanel.stl	1	$3.05	$3.05	Fabricated by JLCPCB	321PA-F NYLON
Enclosure Top	U_Box_V105_General_Alarm_Device_Top.stl	1	$12.4840	$12.48	Fabricated by JLCPCB	321PA-F NYLON
Enclosure Bottom	U_Box_V105_General_Alarm_Device_bottom.stl	1	$13.00	$13.00	Fabricated by JLCPCB	321PA-F NYLON
Enclosure Button	U_Box_V104_General_Alarm_Device_button.stl	1	$1.00	$1.00	Fabricated by JLCPCB	321PA-F NYLON
Enclosure Backpanel	U_Box_V105_General_Alarm_Device_Backpanel.stl	1 NOTE use of this is exclusive of the BackPanel Hanging Compound	$3.34	$3.34	Fabricated by JLCPCB	321PA-F NYLON
Enclosure BackPanel_Hanging	BackPanel_Hanging-CompoundV2.stl	1 NOTE use of this is exclusive of the BackPanel	$5.90	$5.90	Fabricated by JLCPCB	321PA-F NYLON
Enclosure Screws PCB mounting	Choose to fit the enclosure as fabricated	5	$0.10	$0.50	Used #6 sheet metal screws 3/8” long	Steel
Enclosure Screws,	Choose to fit the enclosure as fabricated	4	$0.10	$0.40		Steel
PCB Assembly exclusive of through hole components	BOM_JLCPCB_20230228Modified.xls	1	$13.86	$13.86	Fabricated by JLCPCB	GeneralPurposeAlarmDevicePCB-Placement20230228Modified.xls
U302	LCD Display	1	$4.99	$4.99	https://www.aliexpress.com/item/3256803213374992.html	3256 Display Assembly
Nylon spacers	Spacer_0.0182x0.125 inch	4	$0.13	$0.52	McMasterCarr	94639a702
Screw	4-40x 3/8”	4	$0.12	$0.48	Digikey	36-9901-ND
Nut	Nut_4-40_3/16	4	$0.10	$0.40	Digikey	36-4694-ND
S101	SWITCH_TACTILE_SPST-NO_0.05A_24V	1	$0.13	$0.13	Digikey	450-1804-ND
S601	SWITCH_TACTILE_12mmx12mm_SPST-NO_0.05A_24V	1	$0.55	$0.55	Digikey	SW414-ND
LED White	LED_T1.75_CLEAR_WHITE	6	$0.65	$3.90	Digikey	160-1772-ND
LED Red	LED_T1.75_CLEAR_RED	1	$0.36	$0.36	Digikey	160-1682-ND
RV301	POT 0.375 10K	1	$1.61	$1.61	Digikey	3386P-103LF-ND
9-12V wall power supply	Example Jameco.com 9V 200mA	1	$6.95	$6.95	Jameco.com	100845

These parts sum to $74.43. The printed wiring assemblies were ordered from JLCPCB using two files:
- https://github.com/PubInv/general-purpose-alarm-device/blob/main/Hardware/Manufacturing/BOM_JLCPCB_20230228Modified.xls
- https://github.com/PubInv/general-purpose-alarm-device/blob/main/Hardware/Manufacturing/GeneralPurposeAlarmDevicePCB-Placement20230228Modified.xls

Build instructions

Detailed instructions for the balance of manufacturing on the assemblies received from JLCPCB is in a markdown document, “Manufacturing and Unit Test Documentation, PCB Version 2.0, 20230228” found at:
https://github.com/PubInv/general-purpose-alarm-device/blob/main/Hardware/Documentation/ManufacturingUnitTestTroubleshootingRev2.md
The time to build assemblies was capture in a Github Issue.
See Github Issue #229: The labor to assemble the printed circuit assembly is approximately 40 minutes per unit.
Manufacturing and Unit Test Documentation
PCB Version 2.0
February 28, 2023
This document provides instructions for manufacturing and unit testing of PCB Version 2.0.
1. Manufacturing Instructions
The following steps should be followed to manufacture PCB Version 2.0:
1. Order the required materials from a reputable supplier.
2. Prepare the PCB design for manufacturing.
3. Fabricate the PCB.
4. Assemble the components on the PCB.
5. Test the PCB.
2. Unit Test Instructions
The following steps should be followed to unit test PCB Version 2.0:
1. Connect the PCB to a power supply.
2. Verify that the power supply is providing the correct voltage and current.
3. Inspect the PCB for any visual defects.
4. Test each component on the PCB.
5. Verify that the PCB is functioning correctly.
3. References
1. PCB Design Specification
2. Component Datasheets
3. Test Plan
The following notes were made during the assembly of the JLCPCB order of approximately March 2023. These units will be assigned serial numbers 16 through 30.
Sections below
- Manufacturing test procedure for every unit
  -Assembly Notes and Tips
  - Unit (Design) verification tests on some units
  - Configuration Version 1 Assemblies
- Troubleshooting procedures
Manufacturing test procedure for every unit
Tools Required
- Solder station with appropriate ventilation
- 3/16 Nut Driver
- Number 1 philips screw driver
- Diagonal or other flush cutting hand tool for lead trimming.
- Assembly fixture, detailed below.
Procedure
1. Assemble the unit according to the assembly instructions.
2. Perform the manufacturing test procedure listed below.
3. If the unit passes all tests, it is ready to be shipped.
Manufacturing test procedure
1. Verify that all components are present and correct.
2. Verify that all connections are good.
3. Verify that the unit powers up and runs correctly.
4. Verify that the unit meets all specifications.
Assembly Notes and Tips
- Use a good quality solder and flux.
- Be careful not to overheat the components.
- Trim all leads flush after soldering.
- Use the assembly fixture to help keep the unit aligned during assembly.
Unit (Design) verification tests on some units
- Power-on test
- Functional test
- Performance test
Configuration Version 1 Assemblies
- Use the configuration file provided with the unit.
- Verify that the unit is configured correctly.
Troubleshooting procedures
- If the unit does not power up, check the power supply.
- If the unit does not run correctly, check the connections and the software.
- If the unit does not meet specifications, contact the manufacturer for assistance.

Reference Material

The most recent schematic is for Rev 2 PCB Assemblies is:
https://github.com/PubInv/general-purpose-alarm-device/Hardware/GeneralPurposeAlarmDevicePCB/PDF/Schematic-GeneralPurposeAlarmDevicePCB-V2.2.pdf

Assembly Notes and Tips
- The board will be received from JLCPCB with SMT components placed and some but not all through-hole components.
- The serial number should be written on the PCB assembly at the location indicated. The management of serial numbers is beyond the scope of this document. Each manufacturer must figure something out.
An Assembly Fixture
An assembly assistant/fixture can be made by using a raw PCB with some long #6 screws and nuts to hold at the PCB mounting points.
Please note that these are just some tips and suggestions. The specific assembly process may vary depending on the specific board and components.

In addition to the aforementioned items, four 4-40 screws with washers and 4-40 nuts are required to create a 1/8" spacer for holding the LCD and header for soldering.
Please note that the following changes have been made to the GPAD Rev 2.0 PCB per issue #213:
- Resistor R103 has been removed.
- Factory test for D601 has been updated.
The problem that was identified is that the assemblies built per the BOM_JLCPCB_20230228Modified.xls have R103 fitted with a 1K resistor. This resistor, together with LED D102, loads the SPI_CLK signal and is incompatible with proper operation of the GPAD as an SPI Peripheral from a 3.3V SPI Controller using the level shifting method using the common gate MOSFET.
Please take these changes into account when assembling your GPAD Rev 2.0 PCB.

Solution:
The recommended solution is to remove R103.
Assembly Steps
1. LCD Bezel Grounding
  1. Locate the J1 and J2 solder pads on the LCD module.
  2. Solder them.
2. 16-pin Header
  1. Fit the 16-pin header onto the LCD sub module.
  2. Solder the header pins into the GPAD PCB.
  3. Place four nylon 1/8" spacers at the four corners of the LCD sub module.
  4. Place four 4-40 x 3/8" screws with 4-40 x 3/16" nuts through both boards and torque to 3.4 - 4.8 Inch-Pounds.
3. Buttons and Pot
  1. Place the Reset button, S101, into the PCB from the display side.
  2. Place the Mute button, S401, into the PCB from the display side.
  3. Place the Buzzer, BZ601, into the PCB from the display side. Bending leads may help retain.
  4. Place the Contrast pot, RV301, into the PCB from the display side. Bending leads may help retain.
4. LEDs
  1. Thread the LED leads through the stand-offs.
  2. The longer LED lead is the anode. The cathode has a flat side on the plastic case.
  3. Place the LEDs in to the PCB so that the flat cathode side corresponds to the silk screen marking. Bend the leads to retain the LED into the PCB.
  4. Placing the assembly on the fixture lets you have access to the top and the bottom of the assembly.
  5. Life LED and reflow solder for a flush fit on the PCB.
  6. For each LED, after soldering one lead, holding the LED from below up to the spacer and the PCB and reheat / reflow the solder for a flush fit.
  7. Solder the second lead on the LED.
  8. Trim the excess leads on RV301 and the LEDs.

Assembly Tip: Sharpie Oil-Based Paint Markers can be used to mark polarity on LED standoffs and mark PCB's version, serial number and programming status of the microprocessor.
Congratulations! Electrical assembly is complete.
LCD side components have been installed.
Electrical Tests
Electrical tests are divided into two parts:
1. Unprogrammed measurements made before flash programming the boot loader and other firmware.
2. Programmed measurements made after a boot loader and firmware have been placed into the microcontroller.
Electrical Measurements before programming bootloader and firmware
Measure and record the following electrical parameters by serial number.
Investigate and correct any abnormal measurements before applying power. Remove J102 and J103 and retain them if present. Note where they should be replaced. Start with no connections to the DUT (Device Under Test).
Power Jack Measure resistance to ground at J101 center pin as open or greater than 1Meg ohms.
SPI Interface Measure resistance to ground at J401 pin 5 as open or greater than 1Meg ohms..
Vin net Measure resistance to ground at TP102 as greater than 75K ohms. This net is capacitive and the resistance measured will climb as the multimeter charges the net.
+5V net Measure resistance to ground at TP103 +5 at 1K +/- 5% (950-1050 Ohms).
With a current limited supply set for 12V and maximum of 200 mA, apply power at J101 and note and record the unprogrammed current.
(FYI, when unprogrammed, the first time power up current is normally about 75-80 mA.)
Check that the power LED D105 is lit and is RED.
+5V net Measure the voltage of the +5V Net at TP103.
Please be advised that the current drawn by a programmed DUT that has been powered up and with display backlight on is approximately 61 mA when the reset switch is held down.

Vo Initial Set / LCD Contrast.
With a voltmeter, measure the voltage of the Vo pin of the LCD header to ground. Adjust RV103 for 1.3 V.
See records of measurements of some of the Rev 2 assemblies at this issue: https://github.com/PubInv/general-alarm-device/issues/217 #217

Electrical Test Results Table

- All resistance measured on 20Meg scale except R@5V net measured on 2K scale. Multimeter EMCO Model DMR-3800 unless noted.
- All current measured 200mA scale. Multimeter EMCO Model DMR-3800 unless noted.

Capture:
DUT Serial Number, R@PowerJack, R@SPI Interface, R@Vin net, R@5V net, UnProgramCurrent, Volt@+5 TP103, FullCurrent mA, Vo Volts, Notes

DUT Serial Number	R@PowerJack	R@SPI Interface	R@Vin net	R@5V net	UnProgramCurrent	Volt@+5 TP103	FullCurrent mA	Vo Volts	Notes

Additional Measurements of current with in subcircuits Issue #230
See data from Rev2 build here: https://github.com/PubInv/general-alarm-device/issues/230
Load Firmware
 Note: Loading firmware through the SPI interface is a manufacturing test of the SPI hardware components.
Load Bootloader
Use an Arduino UNO as an ISP (In Circuit Serial Programmer) which will load the boot loader into the DUT.
Cable Connect the ISP UNO to the DUT as follows:

Wiring of the UNO to the DUT.

Place a Jumper on the DUT from D10 to Reset

Missing Jumper to D10 When Load Bootloader

If you forget the above jumper the IDE will give an error something like this:
>Arduino: 1.8.19 (Windows 10), Board: "Arduino Duemilanove or Diecimila, ATmega328P"
>avrdude: Yikes! Invalid device signature.
> Double check connections and try again, or use -F to override
> this check.
>Error while attempting burning of bootloader with out the RJ12 connection.
>This report would have more information with "Show verbose output during compilation" option enabled in File -> Preferences.

Load into the IDE the sketch "ArduinoISP".

To load into the UNO, Select the serial port for the UNO and compile and upload with the "ArduinoISP" by pressing **<Ctrl>U**.

Setup the UNO to burn the boot loader into the GPAD target.
Select the board type (Boot loader type) to "Arduino Duemilanove..." .

Select the Processor type to "ATmega328P" .

Select the programmer type.

In the Arduino IDE, select TOOLS > Burn Bootloader .

Watch the progress bar in the IDE and look for success with the message "Done burning bootloader." in the blue status bar.

IF R103 has not been removed, check that LED D102 is winking with a short on time and longer off time indicating that the boot loader has been loaded. Until any other sketch is loaded this is the expected behaviour of the unit under test.

Load Factory Test Firmware.
Connect a USB cable to the DUT. Note the COM port enumerated in Device Manager Ports(COM&LPT) drop down

In the Arduino IDE, open the new file "GPAD_Factory_Test.ino".
Set the IDE for the COM port of the DUT.
  Using the Arduino IDE, compile and upload to the DUT the "GPAD_Factory_Test.ino"

Watch the progress bar in the IDE and look for success with the message "Done uploading" in the blue status bar.

Open a terminal to the COM port of the DUT and set for appropriate BAUD rate.
Press the reset switch on the DUT and the LCD should display a message. The terminal should display a boot message too. This example is of a RealTerminal connected to the DUT.

Electrical Measurements After Programming Bootloader and Firmware
Measure and record the following electrical parameters by serial number:
Observe the current on the DUT. Press the Mute Switch S601 and the white LEDs D201-D205 should light. The Buzzer will make a sound. Record this full current in the table above.
End of Rev2 Tests as of March 2023

Operation instructions

A video demonstrating [3] of one use of the GPAD (via the serial port) shows not only the usage but also shows the flashing lights and allows you to listen to the audio of escalating urgency.

The GPAD is activated by sending it digital commands over a cable from a different microcontroller. The basic use case is that a condition which needs human attention is detected in whatever manner possible by a sensing microcontroller or computer, called the controller. For example, the sensing device might detect that a temperature is too high, or that a the total flow through an airway in one minute is too low. It than calculates which of five alert levels (not counting the “no alert” condition, which could be considered a sixth) represents the severity of the condition, and a message of up to 80 characters that will inform a person how to remedy the problem. This message is displayed on the back-lit LCD on the GPAD itself. The LCD is dark when the alarm level is silent and back-lit in the other alarm levels.

Conceptually, the sensing device treats the GPAD as an API—an Application Programmers Interface. The six alarm levels are conceptually named:
1. Silent,
2. Informational,
3. Problem,
4. Warning,
5. Critical,
6. Panic.

The exact meaning of the six levels can be defined by the user for the application. Of course, not all six have to be used. However, the effect of the device should be clear: increasing severity leads to increasing sound and light levels produced by the device to obtain the attention of a human being. The GPAD has 5 LEDs, which are normally unlit in the Silent mode. Each increase in alarm level increases the number of LEDs lit. Our intention is that even from across a room a user may comprehend the alarm level at a glance.

Similarly, at increasing urgency levels, the GPAD plays a different rhythmic pattern of high-pitched sirens, which could be described as a “squeal” or a “scream”. This sound currently gets more urgent-sounding in terms of speed, rhythm, and frequency alternation as alarm level increases. Of course, the user is free to change or improve the coding of the GPAD itself, and a programmer will easily see where we implement the “songs''”.

In order to actually send the message, the programmer of the sensing device will likely use our software library. A C/C++ programmer will likely immediately understand the API from its actual declaration in the file alarm_api.h:

#ifndef ALARM_API
#define ALARM_API
#include <Stream.h>

enum AlarmLevel { silent, informational, problem, warning, critical, panic };
// const char *AlarmNames[] = { "OK ","INFO.","PROB.","WARN ","CRIT.","PANIC" };
const int NUM_LEVELS = 6;

const int MAX_MSG_LEN = 80;
const int MAX_BUFFER_SIZE = MAX_MSG_LEN + 1;
typedef struct {
        uint8_t lvl;
        // we will use a null-terminated string!
        char msg[MAX_MSG_LEN+1];
        } AlarmEvent;

int alarm_event(AlarmEvent& event,Stream &serialport);
int alarm(AlarmLevel level,char *str,Stream &serialport);

#endif

The API has two entry points: alarm and alarm_event. The function alarm is simply a convenience function for calling AlarmEvent which constructs an AlarmEvent for the programmer. An AlarmEvent, as stated above, consists of an AlarmLevel and an 80-character (null-terminated) character array holding the message to be displayed on the LCD.

    1. Hardware Connection
Schematics in PDF form can be found in the folder at:
https://github.com/PubInv/general-purpose-alarm-device/tree/main/Hardware/GeneralPurposeAlarmDevicePCB/PDF
You can View Schematics on Line with the links at: https://github.com/PubInv/general-purpose-alarm-device/tree/main/Hardware/GeneralPurposeAlarmDevicePCB
The GPAD requires a 12V power supply via a standard 5.5mm OD 2.1mm pin. It draws less that 1000 mA, even at the highest alarm level.
The GPAD exposes an SPI interface through an RJ12 jack. The SPI pins are mapped the the RJ12 jack as:
SPI Peripheral Pins
The connector is JR12 6P6C and pin 1 is on the left then the connector is held with the contact pins up and facing the observer.
Note the GPAD is a 5V logic level peripheral device. Level shifters must be used when controlled by a 3.3V device.
Pin#	Signal Name	Note
1	nCS	Active low
2	COPI	Input to GPAD
3	GND
4	SCK	Input to GPAD
5	ControllerVcc	5 Volts, an optional way to put power into the GPAD
6	CIPO	Output from GPAD, 5V.

Muting

The GPAD has a mute button. The button silences the sound but does not change the LED lighting. The GPAD stores an internal state as muted or not.

    1. Example Controller Devices

      1. General-purpose Computers

In order to develop and test the GPAD the functioning of the GPAD, it can be driven easily by a serial interface from a laptop computer by simply using a serial monitor. We have designed a trivial interface of commands that the GPAD listens for on the serial interface. Each command begins with one of 4 characters and is terminated by a newline. These are:
1. “a” is the most important command, and is used to set the alarm level. The “a” command is of the “aN”, where N is one of {0,1,2,3,4,5}, specifying increasing alarm levels respectively. These two characters are followed by up to 80 characters defining a message, terminated by an end-of-line (eoln) character.
2. “h” prints a helpful message.
3. “s” means “silence” – mute the GPAD.
4. “u” means “unsilence” or “unmute” – the GPAD is allowed to make noise.

The easiest way to send these serial commands is with a terminal program, or the “serial monitor” feature of the Arduino IDE, which is useful for testing. However, to actually program the GPAD to alert to a recondition, a library such as PySerial that allows a message to be transmitted over the serial port from general purpose programming language.

      1. Microcontrollers
However, an important use case for the GPAD to be programmed by another microcontroller. The SPI interface is particularly useful for controlling a GPAD from another Arduino-class microcontroller. In fact, to make testing as easy as possible, we have software that allows a GPAD peripheral to be programmed by another GPAD acting as a controller. This can be found in the directory GPAD_API_SPI_CONTROLLER (https://github.com/PubInv/general-purpose-alarm-device/tree/main/Firmware/GPAD_API_SPI_CONTROLLER). This code is a useful starting point for anyone programming a GPAD for any given use case, because it exercises the SPI interface and in particular the GPAD Application Programmers Interface.

Validation and characterization

The GPAD is designed to be a part of the Freepsireco ecosystem [2] and serve as the medical alarm in an ICU ward for the PolyVent open-source ventilator. Specifically, it is designed to alert clinicians to conditions which can be electronically computed by the PolyVent microccontroller. These include (but have not yet been coded):
1. A hose disconnect to the patient
2. A loss of pressure
3. Patient-ventilator asynchrony
4. In ability to achieve programmed Peak Inspiratory Pressure, etc.
The GPAD has been demoed in public quite a bit.

After some experimentation to obtain loudness, we built Revision 2 with two different buzzers (Digi-key part numbers 433-WST-1210T-ND and 102-CMI-1295-0585T-ND) which are both rated at 85 dB at 10cm. The LEDs are fairly bright. The basic electrical performance of this simple device seems clear.

However, the true value of an alarm device can only be measured its ability to draw human attention, and then to inform a human being how to take action, and then not to distract them further. There are both simple and complex human/machine interaction performance metrics which we have not begun to measure for the GPAD. A good introduction to the complexity of this field can found in Konkani, Oakley and Bauld (2012).[1]

In the language of the field, the GPAD is not yet a “smart alarm” or an “alarm management system”. It is rather, an annunciation device. (If we called it that, however, non-experts would not understand its intended use.) We hope researchers in this area will find the GPAD a useful programmable component for such such systems and tool for such investigations.

CRediT author statement

Robert L. Read: Conceptualization, Software, Funding Acquisition, Writing – Original Draft.
Lawrence Kincheloe: Supervision, Methodology, Hardware.
(Forrest) Lee Erickson: Software, Validation, Hardware.
Oluseyi 'Seyi' Adeniyi: Software

Lawrence Kincheloe designed the enclosure and mentored the OU capstone projects.
Lee Erickson performed schematic capture, designed the PCBs, wrote the assembly instructions and captured measurements on units during assembly and during some troubleshooting. Lee wrote initial factory test firmware and the factory test procedure which applies the firmware.
Oluseyi 'Seyi' Adeniyi expanded the factory test firmware to include “walking one” like tests for more thorough tests against shorts and bad LCD displays. Seyi also designed the alarm sounds.

Acknowledgments

We would like to thank two senior capstone engineering classes at the University of Oklahoma who made potential extensions to the GPAD. Although they have not yet been used in this project, they have clarified our thinking for future enhancements.

Funding: This work was supported by Public Invention, Austin, TX, a US 501c3 public charity.

References
If relevant, you should include a reference to the original publication of the hardware you customised and a reference to the repository in which your design files are published. Other references can be included, as required; for example, references that put your device in context in the literature. For more information on the reference format in HardwareX please see the Guide for Authors.

[1] A. Konkani, B. Oakley, T. J. Bauld, Reducing hospital noise: a review of medical device alarm management. Biomedical Instrumentation & Technology, 46(6) (2012), 478-487. https://array.aami.org/doi/full/10.2345/0899-8205-46.6.478
[2] R. Read, The Freespireco Project: A complete, free-libre open-source, easily repairable, widely usable, safe and effective respiratory support medical device ecosystem. (2023) https://github.com/PubInv/freespireco.
[3] R. Read, Demo of GPAD Light and Sound. https://youtu.be/F5cf5-Cmwu0. Accessed July 10th, 2023.
[4] Mucco Test Horn literature. https://muccosignal.com/product/warning-horns-with-16-tons/ (accessed 12.08.23)
[5] S.Z. Abidi, V. Suturin, R.L. Read, N. Bechard, 2023, June. A democratized open-source platform for medical device troubleshooting. In 2023 ASEE Annual Conference & Exposition. https://peer.asee.org/a-democratized-open-source-platform-for-medical-device-troubleshooting.
[6] Adafruit Tower Light - Red Yellow Green Alert Light with Buzzer – 12VDC. https://www.adafruit.com/product/2993 (accessed 12.08.2023)
[7] HiLetgo 3D Printer Reprap Smart Controller (Model # 3-01-0889). https://www.amazon.com/HiLetgo-Printer-Controller-Display-Arduino/dp/B07X378P8X. (accessed 12.08.2023).
______________________________________________________________________
Additional Information for authors. (do not include these lines in your submission)
Author manuscript checklist
Is the subject of the submission under an open source license? Are design files in the preferred format for making modifications as defined by the Open Source Hardware definition?

Can the hardware be reproduced with the details provided in the submission?

- Are all relevant design files available on either the Mendeley Data, Open Science Framework, or Zenodo repository? Are they described in the Design Files Summary, and clearly documented? (E.g., descriptive file names, commented code, labeled images, etc.)
  - If in the Open Science Framework, has the repository been registered? Instructions
  - If in Zenodo, is the repository open access and published? Instructions
  - If in Mendeley Data, is the repository published or the sharable link included in the additional information you plan to submit? Instructions

Are visual instructions used when necessary?

Is the utility of the hardware to the scientific community explained clearly? Has a specific scientific application been demonstrated using the hardware?

Is the performance of the hardware adequately demonstrated and characterized?

Are all potential safety concerns addressed?

For more information on the article template consult the Guide to Authors.

Reminder: Before you submit, please delete all
the instructions in this document (the text in italics), including this paragraph.
Thank you!